Gatan Cryoplunge3 General Guidelines

- Open main valve on nitrogen tank and secondary valve that leads to the CP3.

- Connect the liquid nitrogen workstation to the base of the CP3.

- Add the funnel and overflow Styrofoam cup to the left side of the workstation.

- Turn on the CP3 with the switch at the lower right side of the unit.

- Using warm water, thoroughly wet the sponges, insert them into the humidity wand and place the wand in the humidity chamber.

- Prepare the blotters with filter paper and plastic backer using the loading jig.  Insert the blotters into the humidity chamber and set the pressure offset to 0.

- Cool workstation with liquid nitrogen.

- When the workstation has reached equilibrium (~-170 C) you can then pour condensed ethane into the pot.  Monitor the ethane and set the heater appropriately to maintain the ethane just above the melting point.

- After the ethane is added to the pot, use the attached funnel to add necessary liquid nitrogen to the workstation.

- Press ‘START’ to lower the plunge rod.

- Grasp the grid with the tweezers and lock the grid in place by sliding the black clip to the first stop.

- Insert the forceps into the plunge rod until they click in place.  No writing on the forceps will be visible when they are properly in place.

- Raise the rod fully until it clicks in place.

- If necessary, change or rotate the blotters.

- Press ‘RESET’.

- Add sample to the grid through the loading port on the front of the humidity chamber and close the loading port.

- Rotate the plunge rod 90 degrees so the grid is aligned with the blotters.

- Set the timer for the desired blotting time.  For example, when adding 3ul of sample, blot 4 seconds with 2-side blotting and 8 seconds for 1-side blotting.

- Close the safety shield.

- Press ‘START’ to initiate the blotting and plunging sequence.

- Press the blue button on the plunge rod to remove the tweezers.

- Transfer the grid to the grid box.

Proper Clean Up

- Close both valves on the nitrogen supply.  Make sure the valves on the ethane tank are closed.

- Raise the plunge rod to the fully upward position.

- Open the specimen-loading door to air out the humidity chamber.

- Remove the humidity wand from chamber and remove the sponges.

- Press ‘RESET’ several times to bleed the nitrogen gas from the lines.

- Turn off the unit

- Remove the overflow Styrofoam cup and funnel.

- Unplug workstation and dump ethane and nitrogen into ice bucket. Place the workstation on its side to dry.

- Place tweezers, loading jig, humidity wand, sponges and other items used in the storage tray.